## Slide 1: Flora Gateway User Manual
Medical device integration for OR, ICU and bedside workstations
ESM FLORA
Manual 1.0 · gateway release 0.2.4 · October 2026

## Slide 2: What this manual covers
Overview
Where the Gateway sits, its components and data flow
Technology stack
Rust, Python, Kafka, PostgreSQL, Kong, nginx
Using the console
Sign-in, roles, devices, streams, alerts
Canopy integration
Licence, configuration sync, clone, Leaf wizard
Data structures
Relational diagrams for every data store
Reference
Ports, settings, security, updates, troubleshooting

## Slide 3: From device signal to charted observation
Connect
RS-232, HL7 v2 / MLLP, UDP multicast, TCP and HTTP poll / push
Parse
One isolated parser container per device produces Flora observations
Store & serve
Local PostgreSQL store; each Leaf reads its own devices via the data-api
Licence
Device licence, catalogue and parser images come from Haber in Canopy
Watch
Status history for every lane, alerts pushed to Canopy

## Slide 4: Where the Gateway sits in ESM Flora
Notes: Root issues licences and releases; Canopy (with Haber) manages gateways; Leafs read device data from the Gateway on the local network.

## Slide 5: Inside the Gateway: device bytes to Leaf observations
Notes: Controller -> gw.raw.<pod> -> parser -> gw.obs -> collector -> PostgreSQL -> server -> Kong /data-api/<leaf> -> nginx edge -> Leaf.

## Slide 6: Fifteen containers, one Compose project
Container | Role | Port
kafka | Message bus between controllers, parsers and the collector (Kafka 3.9, KRaft) | 127.0.0.1:7414
gateway-db | PostgreSQL 17 store: registry, licence, alerts and observations | 127.0.0.1:7412
serial-controller | RS-232 ingress and egress (frames split by idle time) | 127.0.0.1:7401 admin
socket-controller | TCP/UDP ingress: HL7 MLLP, line, raw, UDP multicast and relay | 7421, 7422, 9500
feeder-controller | Polls device HTTP endpoints | 127.0.0.1:7402 admin
webhook-controller | Receives HTTP POSTs from devices | 7423
parser (one per device) | device-medical-service container, turns frames into observations | internal
collector | Writes observations, measurements and logs to PostgreSQL in batches | 127.0.0.1:7405 admin
publisher | Optionally pushes observations to external HTTP receivers | 127.0.0.1:7406 admin
server | Read API for observations, measurements and device status (Rust/axum) | internal :8080
kong | data-api routes for each Leaf (DB-less Kong 3.9) | 127.0.0.1:7410
gateway-service | Admin console and API, Haber check-in, licence, parser lifecycle, alerts | 127.0.0.1:7400
edge | nginx HTTPS entry point for the console and /data-api/ | 443 (dev 7443)
updater | Applies the release approved in Haber, with health check and rollback | outbound only
simulator (demo) | Synthetic devices for demonstrations and training | 127.0.0.1:7420

## Slide 7: Technology stack
Rust 2021
Controllers, collector, server, publisher: tokio, axum, rdkafka, tokio-serial
Python 3.13
gateway-service (FastAPI 0.118, psycopg 3) and device parsers (aiokafka)
Kafka 3.9.1
KRaft single node, topics per pod, 72 h retention
PostgreSQL 17.6
Registry, licence, alerts and observation time series
Kong 3.9 + nginx 1.27
Per-Leaf data-api routes; HTTPS edge with Flora CA certificates
Docker Compose
One parser container per device; signed, digest-pinned releases

## Slide 8: Protocols and security
Device protocols
HL7 v2 (ORU, IHE PCD-01) over MLLP
Dräger MEDIBUS.X, IACS (M540), Vista CMS link
GE Datex COM 1.2 and DRI
B. Braun BCC (SpaceCom)
JSON push / poll, ASCII key=value
Security
HTTPS edge, certificate from the shared Flora CA
scrypt passwords, HttpOnly 12 h sessions
Admin and operator roles
Bearer gateway key to Haber
Ed25519-signed releases, digest-pinned images

## Slide 9: Using the console
https://<gateway-host> · dev https://localhost:7443

## Slide 10: Sign in and roles
First sign-in
Open https://<gateway-host> and sign in.
New gateway: admin / admin (from GATEWAY_ADMIN_USERNAME / _PASSWORD).
Change it at once: gear → Users → My password.
Sessions last 12 hours; a password change signs out other sessions.
Admin
Everything an operator can do, plus device types, manufacturers, users, the Canopy connection and the site/location.
Operator
Devices (add, edit, enable, disable, remove, restart), controllers (start, stop, restart), streams, acknowledging alerts, Haber sync.

## Slide 11: Four tabs run the gateway
Overview
Station and Canopy status, uptime history, licence, controller layers, live device data, logs
Devices
Installed devices with inventory data; add, edit, enable, remove; per-device Monitor
Streams
Every ingress pod: frames/s, bytes/s, last 200 frames in and out as text or hex
Alerts
Active alerts to acknowledge, alert log, delivery status to Canopy
Gear menu: Device types · Manufacturers · Users · Sign out

## Slide 12: Add a device in four steps
1
Devices → Add device
2
Fill Device ID, type, pod, Leaf and inventory fields
3
Tick Enabled → Save: type, ID and licence are checked, parser starts
4
Monitor: values within a minute; else open Ingress stream
Before you start
The device type exists (from Canopy or a local template)
The licence allows the type and has a free slot
The pod (e.g. socket.icu-monitor) exists in gateway.toml
Device ID is unique: letters, digits, . _ - (max 80)
Changing type, pod or options restarts the parser

## Slide 13: Everyday operations
Restart a parser
Overview → Live device data → Restart
Enable / disable
Devices → Enable/Disable; enabled devices use licence slots
Stop a controller
Overview → Layers → Stop / Start / Restart; stays stopped until started
Look at data
Streams for raw frames · Monitor for parsed values · Overview → Logs
Sync with Canopy
Overview → License → Sync now (check-in immediately)
Remove a device
Devices → Remove; parser stops, stored observations are kept

## Slide 14: Monitoring and alerts
Alert | Category | Severity | Raised when
Canopy connection lost | canopy | warning | No good check-in within 3 × the sync interval
Controller down: <name> | controller | critical | Controller /status unreachable or container stopped
No data from <device> | device | critical | Enabled device sent nothing for 120 s
Gateway service was not running | gateway | warning | Recorded on start when the last sample is more than 2 min old
One status sample per lane every 30 s, kept 7 days
Opens after 2 consecutive down samples; resolves on the first OK
Disabled devices and stopped controllers report stopped, not down
Every change is pushed to the Canopy gateway alert inbox

## Slide 15: Working with Canopy
Licence · configuration sync · clone · Leaf assignment

## Slide 16: Every 20 seconds the gateway checks in

## Slide 17: Connect, license, locate and clone
Connect
Overview → Station → Canopy connection: address, Gateway ID, Gateway key. Saving checks in at once.
Licence
Max devices, allowed types, expiry. Kept while Canopy is offline; revoked or expired stops all parsers.
Location
Set at Canopy → Topology → gateway → Set location…; applied at the next check-in.
Clone
Canopy snapshots each configuration; clone it to a new gateway, which receives it on first check-in.

## Slide 18: Assign devices to a Leaf with the gateway wizard
1
Test the data-api URL the gateway reported
2
Choose devices for this Leaf
3
Assignment travels via Canopy, applies at next check-in
4
Verify values from the last 60 s, save the source
On the Leaf: Ward settings → Gateway wizard (permission config.manage). The Leaf then reads https://<gateway>/data-api/… filtered by its leaf_id.

## Slide 19: Supported device types
Code | Device | Controller | Parser
hl7-patient-monitor | Patient monitor (HL7 v2 / IHE PCD-01) | socket | hl7v2
hidro-hl7-monitor | Hidro shared HL7 monitor | socket | hidro_hl7
draeger-iacs-m540 | Dräger M540 (IACS UDP) | socket | iacs_m540
draeger-vista-120s | Dräger Vista 120 S (CMS link, TCP) | socket | draeger_vista_s
draeger-medibus | Dräger MEDIBUS.X (RS-232) | serial | medibus
ge-carestation-750 | GE Carestation 750 (COM 1.2) | serial | ge_carestation
ge-aisys-cs2 | GE Aisys CS2 (COM 1.2) | serial | ge_carestation
ge-bx50-dri | GE Bx50 / B650 (DRI) | serial | ge_dri
ge-s5-dri | GE S/5 (DRI) | serial | ge_dri
ge-aisys-dri | GE Aisys (legacy DRI) | serial | ge_dri
bbraun-space-bcc | B. Braun SpaceCom (BCC over TCP) | socket | bbraun_bcc
json-anesthesia-machine | Anaesthesia machine (JSON push) | webhook | json_fields
http-ventilator | ICU ventilator (HTTP poll) | feeder | json_fields
serial-temp-module | Temperature / CVP module (RS-232, polled) | serial | ascii_kv
acme-t200 | Acme T-200 temperature / CVP (RS-232) | serial | acme_t200

## Slide 20: Data structures
Relational diagrams · solid = foreign key · dashed = logical link

## Slide 21: Gateway DB · device registry

## Slide 22: Gateway DB · licence, site, Canopy link, users

## Slide 23: Gateway DB · observations (time series)

## Slide 24: Gateway DB · monitoring, alerts, operations

## Slide 25: Canopy DB · gateway configuration and alerts

## Slide 26: Haber SQLite · gateway registry and releases

## Slide 27: Kafka topics · message envelopes

## Slide 28: Network and ports
Port | Service | From
443 (dev 7443) | HTTPS edge: admin console and /data-api/ | LAN
7400 | gateway-service admin (HTTP) | localhost
7401–7406 | Controller, collector and publisher /status | localhost
7410 | Kong data-api (HTTP) | localhost
7412 / 7414 | PostgreSQL / Kafka | localhost
7421 / 7422 | HL7 MLLP OR / ICU monitors | device LAN
7423 | Device webhook (HTTP POST) | device LAN
9500 (UDP 9610) | Dräger Vista 120 S CMS link (beacon) | device LAN
2050 UDP · 2000 | Dräger M540 data · discovery multicast 224.127.1.254 | device LAN
2051 | M540 relay (udp_relay framing) | device LAN
4001 (outbound) | B. Braun BCC SpaceCom | device LAN
7420 | Simulator (demo profile only) | localhost
Canopy 443 (dev 9443) | Haber check-in at /api/haber/ (outbound) | hospital

## Slide 29: Key settings in flora-gateway/.env
Variable | Default | Meaning
FLORA_GATEWAY_ID | gateway-hospital-01-a | Gateway identity, stamped on every frame
HABER_URL | https://host.docker.internal:9443 | Canopy / Haber address
HABER_GATEWAY_TOKEN | (demo token) | Gateway key sent as Bearer token
GATEWAY_SYNC_INTERVAL_SEC | 20 | Check-in interval in seconds
GATEWAY_LICENSE_REQUIRED | true | Without a valid licence, parsers do not run
GATEWAY_DATA_API_URL | http://host.docker.internal:7410/data-api | data-api URL reported to Canopy for the Leaf wizard
GATEWAY_ADMIN_USERNAME / _PASSWORD | admin / admin | First admin account (only when no users exist)
GATEWAY_SESSION_HOURS | 12 | Console session lifetime
GATEWAY_ADMIN_KEY | (empty) | Optional x-gateway-key header for scripted write calls
FLORA_GATEWAY_CONFIG_FILE | gateway.toml | Controller and pod configuration
FLORA_GATEWAY_SEED_FILE | instances.json | First-start device seed (ignored once Canopy config applied)
FLORA_GATEWAY_DB_PASSWORD | flora-gateway-local-only | PostgreSQL password — change in production

## Slide 30: Releases arrive signed and approved
Root signs
Haber mirrors
Hospital approves
Updater applies
Health check / rollback
Images are pinned by sha256 digest and verified against the Root key
Health checks within 180 s; automatic rollback if unhealthy
Parser images upgrade one device per check-in cycle
Optional maintenance window: FLORA_UPDATER_WINDOW=HH:MM-HH:MM

## Slide 31: Troubleshooting
Symptom | What to do
"unknown device type …; sync from Haber first" | The type is not in the catalogue yet. Check the Canopy connection and the licence, then Overview → License → Sync now.
403 "device license invalid / not licensed / device limit reached" | Licence problem: expired, revoked, type not allowed or max_devices reached. Ask the Canopy administrator; Root issues licences per hospital.
Canopy shows 503 "no verified license bundle" | Canopy has no verified bundle from Flora Root yet: a Canopy-side issue.
Alert "No data from <device>" | Open Streams: are frames arriving on the pod, are polls going out? Check cable, baud/parity, port, source IP and that the controller is running.
Controller shows Stopped | It stays stopped across restarts until you press Start (or run docker compose up).
COM port not visible (Windows) | Docker Desktop cannot pass COM ports. Run serial/socket controllers natively and point SERIAL_ADMIN_URL / SOCKET_ADMIN_URL to them.
Dräger M540 silent | Multicast does not cross subnets or Docker bridges: use tools/iacs-relay/relay.py with udp_relay framing.
Edits to instances.json have no effect | It is only a first-start seed. Edit devices in the console.

## Slide 32: Quick reference
Console https://<gateway> · data-api /data-api/<leaf-id>/api/observations · full manual and diagrams in docs/gateway-manual/